Test Results

Respirators help protect against certain particulate contaminants but do not eliminate exposure to or the risk of contracting any disease or infection. Misuse of a respirator can result in exposures that can result in sickness or death.

The test results below indicate the performance of the filtration material specific to the respirator tested. The materials were tested using specialized equipment and protocols, similar but not equivalent to the NIOSH standard procedure TEB-APR-STP-0059; reported filtration efficiencies may differ accordingly. These results do not guarantee that all respirators produced by a manufacturer will have a similar level of filtration efficiency or a similar level of resistance to fluid penetration. These results do not imply a fitness for a particular use or compliance with any federal or international standards. These respirators were not tested in an accredited National Institute for Occupational Safety and Health (NIOSH) testing laboratory, and thus are not NIOSH approved or certified. We have worked with the Advanced Functional Fabrics of America, Inc. (AFFOA) to obtain respirator performance characterization. AFFOA coordinated testing at MIT Department of Chemical Engineering and MIT Lincoln Laboratory where comparable conclusions about mask quality were obtained. The Filtration efficiency measurements listed below were taken by the Massachusetts Institute of Technology. The resistance to fluid penetration measurements below were taken by MIT Lincoln Laboratory.

According to ASTM F 1862, surgical masks are tested on a pass/fail basis at three velocities corresponding to the range of human blood pressure (80, 120, 160 mm Hg). Fluid resistance may be claimed if the device passes ASTM F1862 at the corresponding level. Surgical masks that show passing results at higher velocities are more fluid resistant. The test results below indicate the performance of the specific mask tested.

POWECOM

Test date – 4/18/20 - Filtration efficiency: 90.6% Test date – 4/17/20 – Filtration efficiency: 92 %

Visit CDC website for additional information on this  product

|  | Jiangsu Nanfang Medical CO. LTD Test date – 4/17/20 – Filtration efficiency: 90.3% Test Date – 4/24/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through |
| --- | --- |
|  | Foshan Wei hui Labor Protection Products Co., Ltd Test date – 4/17/20 – Filtration efficiency: 84.2% Test Date – 4/24/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through, but partial layer penetration along seams |
|  | Unlabeled – plastic bag with KN95 on the mask Test date – 4/17/20 – Filtration efficiency: 28.1% |
|  | KN-95 Disposable Face Mask Test date – 4/16/20 – Filtration efficiency: 52.3% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through, but partial layer penetration along seams |
|  | SHUYIAN Test date – 4/16/20 – Filtration efficiency: 87.6% |

|  | PinzTec Test date – 4/16/20 – Filtration efficiency: 83.6% |
| --- | --- |
|  | SDI Folding Anti-Particulate respirator Test date – 4/16/20 – Filtration efficiency: 91.4% |
|  | ESOUND Med Test date – 5/18/2020 – Filtration Efficiency: 96.4% Test Date – 4/29/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through |
|  | Drager Test date – 4/24/2020 – Filtration efficiency: 95% Test Date – 4/29/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through |
|  | AOXING Test date – 4/24/2020 – Filtration efficiency: 87.8% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through, but partial layer penetration along seams |

|  | Dasheng Model DTC3X Test Date – 4/24/2020 – Filtration efficiency: 88.7% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through Visit CDC website for additional information on this product |
| --- | --- |
|  | Americares Emergency Response BYD CARE Test Date – 4/24/2020 – Filtration efficiency: 89% Test Date – 4/29/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through |
|  | KN95 (FFP2) Test Date – 4/24/2020 – Filtration efficiency: 70% Test Date – 4/29/2020 – Penetration Resistance Test: Failed |
|  | GONGRENCHENGPING Respiratory Face Mask Test Date – 4/24/2020 – Filtration efficiency: 85.9% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); but partial layer penetration along stitches |
|  | LVJIANGNAN Test Date – 4/24/2020 – Filtration efficiency: 91.4% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through |

|  | DASHENG MODEL DTC3W Test Date – 4/26/2020 – Filtration efficiency: 91.8% Test Date – 4/27/2020 – Penetration Resistance Test: Failed |
| --- | --- |
|  | Saiertech Test Date – 4/26/2020 – Filtration efficiency: 85.9% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through, but partial layer penetration along seams |
|  | OANY Test Date – 4/26/2020 – Filtration Efficiency: 62.3% Test Date – 4/29/2020 – Penetration Resistance Test: Failed |
|  | Gerber Outerwear (White Packaging) Test Date – 5/27/2020 – Filtration efficiency: 83.4% Test Date – 4/29/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); no soak through |
|  | PURVIGOR Test Date – 4/27/2020 – Filtration efficiency: 73.4% Test Date – 4/29/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); but partial layer penetration along stitches Visit CDC website for additional information on this product |

|  | JCH Industrial, Hong Kong Test Date – 4/27/2020 – Filtration efficiency: 86.7% Test Date – 4/29/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); but partial layer penetration along stitches |
| --- | --- |
|  | Lamdown Test Date – 4/27/2020 – Filtration efficiency: 87.4% Test Date – 4/27/2020 – Penetration Resistance Test: Passed High Standard (Level 3, 160mmHg); but partial layer penetration along stitches |
|  | N/A Test Date – 4/27/2020 – Filtration efficiency: 62.6% Test Date – 4/29/2020 – Penetration Resistance Test: Failed |
|  | Yiwu Hongbiao Garments Co. Ltd Test Date – 4/29/2020 – Filtration Efficiency: 91.1% Test Date – 4/24/2020 – Liquid Penetration Test: Passed |
|  | 3M KN-95 Campco, Inc. Test Date – 4/29/2020 – Filtration Efficiency: 95.6% Test Date – 4/24/2020 – Liquid Penetration Test: Passed |
|  | Makrite Test Date – 4/26/20 – Filtration Effeciency: 91.6% Test Date – 4/25/2020 – Liquid Penetration Test: Passed |
|  | Safe Secure Test Date – 4/29/2020 - Filtration Efficiency: 89.8% Test Date – 4/24/2020 – Liquid Penetration Test: Failed, MP 1 /2 soak through in main body when pressed with swab |
|  | GB2626-2006 KN95 Test Date – 4/29/2020 - Filtration Efficiency: 78.9% Test Date – 4/25/2020 – Liquid Penetration Test: Passed |
|  | PM 2.5, 5-pack Test Date – 4/29/2020 - Filtration Efficiency: 73.20% Test Date – 4/25/2020 – Liquid Penetration Test: Passed |
|  | HG, 5-pack Test Date – 4/29/2020 - Filtration Efficiency:51.10% Test Date – 4/25/2020 – Liquid Penetration Test: Passed |
|  | ONE MIX Test Date – 4/29/2020 - Filtration Efficiency:85.3% Test Date – 5/18/2020 – Filtration Efficiency: 85.9% Test Date – 4/27/2020 – Liquid Penetration Test: Passed |
|  | Huran Test Date – 4/29/2020 - Filtration Efficiency:93.00% Test Date – 4/27/2020 – Liquid Penetration Test: Passed |
|  | Ace Neale Test Date – 4/29/2020 - Filtration Efficiency: 92.0% Test Date – 4/28/2020 – Liquid Penetration Test: Passed |
|  | Guan Hua Item 9801 Test Date – 4/29/2020 - Filtration Efficiency:84.7% |
|  | Linxia Test Date – 4/29/2020 - Filtration Efficiency:89.2% |

|  | KANGERDA Test Date – 4/29/2020 – Filtration Efficiency: 70.5% Test Date – 4/24/2020 – Liquid Penetration Test: MP, HP ½ soak through center seam when pressed with swab |
| --- | --- |
|  | DROMEX (Campo Inc.) Test Date – 4/30/2020 – Filtration Efficiency: 80.9% Test Date – 4/24/2020 – Liquid Penetration Test: Passed |
|  | S. Sansgo Test Date – 4/30/2020 – Filtration Efficiency: 90.0% |
|  | ZHIYI Test Date – 5/5/2020 – Filtration Efficiency: 82.4 % Test Date – 5/4/2020 – Liquid Penetration Test: Liquid Penetration Test: Failed LP 1 /2 oak through at center seam HP 1 /2 almost soak through at center seam |
|  | Chun Lam Group Test Date – 5/5/2020 – Filtration Efficiency: 86.0% Test Date – 5/5/2020 – Liquid Penetration Test: Passed |
|  | DR. MFYAN Test Date – 5/5/2020 – Filtration Efficiency: 45.5% % Test Date – 5/5/2020 – Liquid Penetration Test: Liquid Penetration Test: Failed partial penetration through outer layers |
|  | ALLMAN Test Date – 5/5/2020 – Filtration Efficiency: 92.4 % Test Date – 5/5/2020 – Liquid Penetration Test: Passed |
|  | Doctor Hua (Jiayuan Technology) Test Date – 5/5/2020 – Filtration Efficiency: 63.4 % Test Date – 5/5/2020 – Liquid Penetration Test: Passed |
|  | Shenzhen Zhiyou Test Date – 5/5/2020 – Filtration Efficiency: 61.8% Test Date – 5/5/2020 – Liquid Penetration Test: Passed |
|  | Shanghai Huaxiang Woolen Dressing Huaxiang Woolen Test Date – 5/5/2020 – Filtration Efficiency: 84.1 % Test Date – 5/5/2020 – Liquid Penetration Test: Passed |
|  | China Southern Test Date – 5/5/2020 – Filtration Efficiency: 72.8 % Test Date – 4/22/2020 – Liquid Penetration Test: Liquid Penetration Test: HP 1/2 soak through in main body of filter when pressed with swab |
|  | Ocean State Donation Test Date – 5/5/2020 – Filtration Efficiency: 86.0 % |
|  | FFP2 NK International (King Year Printing & Packaging Co. Ltd.) Test Date – 5/5/2020 – Filtration Efficiency: 75.4% Test Date – 5/26/2020- Filtration Efficiency: 81.7% Test Date - 4/22/2020- Penetration Resistance: Passed; no soak through, but some penetration along the stitching |
|  | Model:YH006 (Joint China Ltd.) Test Date – 5/5/2020 – Filtration Efficiency: 81.6% Test Date – 5/5/2020 – Liquid Penetration Test: Passed |
|  | BIWEIKANG Test Date: 5/8/2020 – Filtration Efficiency: 85.5% |
|  | Easewell (Easepal Ind. Co. Ltd.) Test Date: 5/8/2020 – Filtration Efficiency: 81.7% Test Date – 5/4/2020 – Liquid Penetration Test: Passed |
|  | Makrite (SEKURA) Test Date: 5/8/2020 – Filtration Efficiency: 92.6% Test Date – 5/8/2020 – Liquid Penetration Test: Passed |
|  | EATA Test Date: 5/8/2020 – Filtration Efficiency: 73.2% Test Date – 5/4/2020 – Liquid Penetration Test: Passed |
|  | Uniair (SAN HUEI United Co. Ltd) Test Date: 5/8/2020 – Filtration Efficiency: 88.9% Test Date – 5/8/2020 – Liquid Penetration Test: Passed |
|  | Xier Test Date: 5/12/2020 – Filtration Efficiency: 62.0% Test Date – 5/12/2020 – Liquid Penetration Test: Passed |
|  | Spectrum 50-pack Test Date: 5/12/2020 – Filtration Efficiency: 72.5% Test Date – 5/11/2020 – Liquid Penetration Test: Passed |
|  | San Huei United Co., Ltd. SH9550 Test Date: 5/12/2020 – Filtration Efficiency: 86.4% Test Date – 5/12/2020 – Liquid Penetration Test: Passed |
|  | In Jeolly Test Date: 5/12/2020 – Filtration Efficiency: 83.4% Test Date – 5/12/2020 – Liquid Penetration Test: Passed |
|  | Fujian Meinkind Baby Products, Ltd Test Date: 5/12/2020 – Filtration Efficiency: 71.1% Test Date – 5/12/2020 – Liquid Penetration Test: Passed |
|  | 50 pack donation Test Date: 5/12/2020 – Filtration Efficiency: 91.1% |
|  | Aumacom Test Date: 5/12/2020 – Filtration Efficiency: 48.5% |
|  | XAM-Med Test Date: 5/12/2020 – Filtration Efficiency: 85.7% |
|  | ANHUIYAXIA Anhui Yaxin Plastic Test Date: 5/20/2020 – Filtration Efficiency:85.0% Test Date – 5/20/2020 – Liquid Penetration Test: Passed |
|  | MINGCAI Test Date: 5/20/2020 – Filtration Efficiency: 61.0% Test Date – 5/20/2020 – Liquid Penetration Test: Passed |
|  | Benehal Suzhou Sanical Benehal Test Date: 5/20/2020 – Filtration Efficiency:85.0% Test Date – 5/23/2020 – Liquid Penetration Test: Failed, soak through at center seam |
|  | China Etech Test Date: 5/20/2020 – Small Font Filtration Efficiency:86.1% Large Font Filtration Efficiency: 87.0% |
|  | AM – AiMei Test Date: 5/20/2020 – Filtration Efficiency: 81.6% |
|  | Gansu Newshanyou Test Date: 5/20/2020 – Filtration Efficiency: 59.7% Test Date – 5/20/2020 – Liquid Penetration Test: Passed |
|  | HY - Hunan Huanyu Test Date: 5/20/2020 – Filtration Efficiency: 54.5% Test Date – 5/20/2020 – Liquid Penetration Test: Passed |
|  | ESound Med, “N-95” Test Date: 5/29/2020- Filtration Efficiency: 97.19% |
|  | Donation, 30 pack Test Date: 4/22/2020- Filtration Efficiency: 92.3% |
|  | Sky Screen, Box of 20 Test Date: 4/22/2020- Filtration Efficiency: 98.8% Test Date – 4/22/2020 – Liquid Penetration Test: Passed |
|  | Donation, 2-pack Test Date: 5/11/2020- Filtration Efficiency: 77% |
|  | Guo Xiang, 3 pack Test Date: 5/11/2020- Filtration Efficiency: 82.1% Test Date – 5/12/2020 – Liquid Penetration Test: Failed soak through at center seam |
|  | Hong Jie, Box of 10 Test Date: 5/11/2020- Filtration Efficiency: 81.3% Test Date – 5/11/2020 – Liquid Penetration Test: Passed |
|  | King Fa Test Date: 5/11/2020- Filtration Efficiency: 87.4% |
| Spectrum 1-pack | Spectrum, 1-pack Test Date: 5/11/2020- Filtration Efficiency: 76.6% Test Date – 5/11/2020 – Liquid Penetration Test: Passed |
|  | Face Code, package of 10 Test Date: 5/23/2020, Filtration Efficiency: 86.3% Test Date – 5/23/2020 – Liquid Penetration Test: Passed, but front is absorbent |
|  | Fujian Pageone, Package of 5 Test Date: 5/23/2020, Filtration Efficiency, 87.5%-99% Test Date – 5/23/2020 – Liquid Penetration Test: Passed |
|  | Fujian Pageone, Package of 5, Marked “Non-Medical” Test Date@ UML: 1/10/2021: Filtration Efficiency 93.26%-96.89% Liquid Penetration Test: Passed Test Date @ MIT: 1/10/2021 Filtration Efficiency: 98.3%-99.5% |
|  | Shenzen YiFan, Package of 5 Test Date: 5/23/2020, Filtration Efficiency: 55.2% Test Date – 5/23/2020 – Liquid Penetration Test: Passed |
|  | Shenzen YiFan, Package of 5 (Version B, different packaging) Test Date: 6/23/2020, Filtration Efficiency: 86.4% Test Date – 6/23/2020 – Liquid Penetration Test: Passed |
|  | Etech, Box of 50 Test Date: 5/18/2020, Filtration Efficiency: 86.1% Test Date – 5/18/2020 – Liquid Penetration Test: Passed |
|  | Zhong Yi, Box of 20 Test Date: 5/26/2020, Filtration Efficiency: 98.9% |
|  | Reis, Pack of 5 Test Date: 5/18/2020, Filtration Efficiency: 82.1% Test Date – 5/18/2020 – Liquid Penetration Test: Passed |
|  | Dasheng (green box, Model DTC3X), Box of 20 Test Date: 5/26/2020, Filtration Efficiency: 98.9% |
|  | Allta, Box of 20 Test Date: 5/26/2020, Filtration Efficiency: 88.4% |
|  | Seumex, Box of 10 Test Date: 5/26/2020, Filtration Efficiency: 34.4% |
|  | ZheJaing, Box of 50/pack of 10 Test Date: 5/26/2020, Filtration Efficiency: 29.9% |
|  | Wei Hui, Version 2 (Different Packaging), Box of 50 Test Date: 6/1/2020, Filtration Efficiency: 42.6% |
|  | EsMoo, Box of 50 Test Date: 6/1/2020, Filtration Efficiency: 99.5% |
|  | King Year, Box of 20 Test Date: 6/1/2020, Filtration Efficiency: 74.5% |
|  | Sinopharm. Package of 1 Test Date: 6/1/2020, Filtration Efficiency: 96.7% |
|  | Dongguan, Box of 20 Test Date: 6/1/2020, Filtration Efficiency, 97.7% |
|  | Gerber (Green Packaging), Package of 10 Test Date, 5/26/2020, Filtration Efficiency: 38.3% |
|  | Daddy Baby (Blue Box), Box of 50 Test Date: 6/15/2020, Filtration Efficiency: 65.3% Test Date – 6/15/2020 – Liquid Penetration Test: Passed |
|  | Daddy Baby (Green Box), Box of 50 Test Date: 6/15/2020, Filtration Efficiency: 94.1% Test Date – 6/15/2020 – Liquid Penetration Test: Passed |
|  | Scott Wang Brand, Bag of 20 Test Date: 6/15/2020, Filtration Efficiency: 24.9% Test Date – 6/152020 – Liquid Penetration Test: Passed |
|  | GearMesh, Box of 20 Test Date: 6/15/2020, Filtration Efficiency: 78.6% Test Date – 6/15/2020 – Liquid Penetration Test: Passed |
|  | Tomson Newt, Package of 2 Test Date: 6/15/2020, Filtration Efficiency: 96.2% Test Date – 6/15/2020 – Liquid Penetration Test: Passed |
|  | Zateny, Box of 10 Test Date: 6/15/2020, Filtration Efficiency: 98.0% Test Date – 6/15/2020 – Liquid Penetration Test: Passed |
|  | Fido, Box of 20 Test Date: 6/23/2020, Filtration Efficiency: 99.6% Test Date – 6/23/2020 – Liquid Penetration Test: Passed |
|  | Zhuyi Medical Version A, Box of 50 Test Date: 4/24/2020, Filtration Efficiency: 49.5% Test Date – 4/24/2020 – Liquid Penetration Test: Passed |
|  | Zhuyi Medical, Version B (Different Packaging), Box of 50 Test Date: 6/29/2020, Filtration Efficiency: 74.0% Test Date – 6/29/2020 – Liquid Penetration Test: Passed |
|  | Zhongshan First Mountain Electrical Appliance Company Limited- F002 Test Date: 10/1/2020: Filtration Efficiency: 85.5% Test Date: 10/1/2020: Liquid Penetration Test: Passed |

This document will be updated as additional masks and samples are tested